## Slide 1: Run2-TDC Development and Test
Su-Yin Grass Wang, Jin-Yuan Wu, David Christian Da-Shung Su and Wen-Chen Chang

## Slide 2: Outline
Improvements of Run2 TDC
Run2 TDC design
Deadtime estimation
Production test (with CODA)
  Bin-Width Plot
  Test # 1~5 (Cross-Talk, Resolution, Multi-Hits for Buffer 1~3, remove multiple hit function test)
Time line

## Slide 3: Improvements of Run2-TDC
Zero suppression is naturally applied
  Based on the multi-sampling principle, the TDC card records hits only. Zero suppression is intrinsic in the design.
Adjustable time window
  The time window can be adjusted to fit all E906 detectors
Multiple events per interrupt (2 ~ 32, user defined)
Multiple hits elimination - keep only first hit in user select time region.
Scalar - keep counting the number of hits in each channel

## Slide 4: Run2 TDC Design
Input signals propagate in a 9-bin delay line.
There is a free running 250MHz clock inside the TDC.
Input transitions are recorded every 4 ns.
4ns / 9 bins~0.44ns/bin

## Slide 5: Run2 TDC Design
First layer buffer
second layer buffer
Event buffer
The TDC records the time of every hit in each channel. A maximum of 4 hits per channel can be temporarily stored in the first layer buffer.
Every 4 clock cycles at 62.5 MHz, one hit in first layer buffer of each channel is transferred to the second layer and cleared. Data from a group of 4 channels are written into one of the 2 (or 4,8) circular buffers inside FPGA, in await for a trigger.
When a trigger arrives, the pointer of writing is shifted to the next circular buffer and the hits in the circular buffer are copied to the event buffer.

## Slide 6: Multiple Hits Elimination
TFQ[0] (Data valid)
Reg1_0_Q[0]
DFN1_1_Q
(DFN1_2Off if DD[6]=1)
DFN1_0_Q
(DFN1_2On if DD[6]=1)
DFN1_2_Q
3
3
2
1
0
3F
3E
3D
X
Q[5:0]
3C
3
DFN1_3_Q
Hits from Reg1_0_Q[0] are blocked in this region.
DD is the register controls multiple hits elimination.
DD[7]
  =1: Eliminating Multiple Hits
DD[6]
  =1: Updating Mode
  =0: Non-updating Mode
DD[5:0]
  0-63: Number of cycles of multiple hits elimination. 4ns*(4+DD[5:0])
  (updating mode 0-31 only)

## Slide 7: Multiple Hits Elimination
In register map : 0x00c
  BLKS(26-24) |DeMultiDD(23-16) |CR_Reset(15) | CR_Clear(0)
  DeMultiDD (5-0) length of Inhibit : 0-63: Number of cycles of multiple hits elimination. 4ns*(4+DD[5:0]) (16 – 272 ns)
  (updating mode 0-31 only)
  DeMultiDD (6) update mode/NON –update mode
  DeMultiDD (7) Elimination on/off
Raw data
Inhibit
TDC recorded

## Slide 8: Scalar
TC[2]
DD is the register controls multiple hits elimination.
of multiple hits elimination. 4ns*(4+DD[5:0])
62
63
0
1
2
3
4
5
TC[8..3]
62
63
0
1
2
3
4
5
TCQ[2]
TCQ[8..3]
SE[0]
SE[1]
CH0
1
2
3
4
5
QS[7..0]
SEL=TC[2]
62
63
0
1
2
3
4
5
ADDRA[5..0]
=TC[8..3]
(for read)
62
63
0
1
2
3
4
5
ADDRA[5..0]
(for write)
Write
(inv. RWA)
=TC[2]

## Slide 9: Run2 TDC Design
TDC bin width | ~0.44 ns
Minimum width of signal | 4 ns
Maximum number of hits in 64 ns | 4 hits
Width of accepted window | Adjustable: 4 – 2048 ns in a 4-ns step
Edge detection | Adjustable: leading edge only or both edges.
Maximum number of hits per trigger | Adjustable: 32-1024
Maximum numbers of events per IRQ | Adjustable: 2-32
Scalar buffer to record # of hits each channel | 8
Multiple Hits Elimination | Adjustable: Disable time window (16 -272)

## Slide 10: Register Map
0x000 version number (e9061212)
0x004 CSR
0x008 test area =0xaaaaaaaa
0x00c
  BLKS(26-24) |DeMultiDD(23-16) |CR_Reset(15) | CR_Clear(0)
0x010 0xe906e0## = # of events
0x014
  Time Window On(31)
  UP_LIMIT(25-16) | LOW_LIMIT(9-0)
0x030 –Time window Check
0x1000~0x1ffc = event buffer (1k internal)
0x2000~0x2ffc = test area
D31-D16 | D11 | D08 | D7-5 | D2-0
EN4X (CSR) | Trig Enable | BothEdge (Edge selection) | SEL(READ/WRITE addr) | M (buffer size)

## Slide 11: Configuration
Event table
  Adjustable
    Time window width
    CIP time
CSR [7:5] | Number of Circular Buffers | Circular Buffer Size (MB) | Buffer Timing Window | CIP Time
000 | 8 | 32 | 512 ns (-256 ns) | 32 x 16 x 16 ns = 8 us
001 | 4 | 64 | 1024 ns (-256 ns) | 64 x 16 x 16 ns = 16 us
011 | 2 | 128 | 2048 ns (-256 ns) | 128 x 16 x 16 ns = 32 us
CSR [2:0] | Number of events | Max. number of words/event
5 | 32 | 32
4 | 16 | 64
3 | 8 | 128
2 | 4 | 256
1 | 2 | 512
0 | 1 | 1024
Event Buffer
  Adjustable
    buffer size
    number of buffer
Notes: These buffer size corresponding to hit accepting window of 512, 1024 and 2048 ns. Due to effect of the RegPipe4a, the 100% hit accepting windows are 256 ns narrower.
The 256 ns here is the part close to the trigger of TDC. Since our trigger is delayed by 700us, we may not care.

## Slide 12: DATA format
D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D0
Trig [15:4] (trigger time) |  |  |  |  |  |  |  |  |  |  |  | Fine time |  |  | 
D31 | D30 | D29 | D28 | D27 | D26 | D25 | D24 | D23 | D22 | D21 | D20 | D19 | D18 | D17 | D16
1 | CNT[10:0] (number of hits) (could indicate truncated data readout ) |  |  |  |  |  |  |  |  |  |  | 0 | 0 | 0 | 1
D31 | D30 | D29 | D28 | D27 | D26 | D25 | D24 | D23 | D22 | D21 | D20 | D19 | D18 | D17 | D16
Cable: 4,5,6,7 |  |  |  | Channel number: 0-15 |  |  |  | TriggerID (for multiTrigger setup) |  |  |  | 0 | 0 | 0 | positive or negative edge
Header
Data
D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D0
TDC [15:4] (data time) |  |  |  |  |  |  |  |  |  |  |  | Fine time |  |  | 
Error check or checksum?
Notes: MSB = most significant bit

## Slide 13: CODA raw data
Ignore every 0xe906e906 in DATA
header
Data
boardID
TDC setup= 0xffff0000&Reg(0x00c) +0x0000ffff&Reg(0x004)
FEE type
End of TDC data

## Slide 14: Data transfer speed
A Gigabit Ethernet support at 1 Gbps (1000 Mbps =125 MBps)
VME backplane speed for BLT32 = 25 MB/s (25 B/us) = (10 us to readout each TDC)
In comparison with TDC readout speed, the internet speed won’t be the problem.

## Slide 15: Minimum time to operate each TDC event
10 us for each IRQ
1.950+ +1.230+1.920= 5.1 us
For each 32bit data point (each hit)= 234.17ns
(~17Byte/us) the TDC data transfer will be the main source of dead time
Trigger Enable
Initialize TDC Header
Check Tdc csr
Trigger Disable
Read Header File
If there are data in buffer, read data

## Slide 16: Real case: number of hits per ROC
In the final run (2173) of Run1
In the worst scenario, there are 750 hits per ROC per event .
To read all hits in the ROC, we need 750 *234.17ns =175.6 us
Suppose that there are 7 TDCs in that ROC, the dead time will be 175.6 (234ns *750 hits) + 7*5.1 (5.1us* 7 TDCs) +10 (IRQ)=221 us
That is much better than 680 us in the first run.

## Slide 17: Roc 16 is usually the one with the largest num. of hits
In average, number of hits per ROC per event is less than 200.
46.9 (234ns *200 hits) +35.5 (5.1us* 7 TDCs) +10 (IRQ) = 92.4 us

## Slide 18: Deadtime estimation with Beam structure
Bin size =137 us
It looks that we won’t lose too much event when dead time is short than 137us.
Dead time was 680 us in Run1
If max hits per ROC is more than 391(=(137-35.5-10)/0.234), the dead time will be greater than 137us.
In Run 2173 there are 4852/26225 = 18.5% events with max hits more than 391 on ROC16.
Before Inhibit
After Inhibit
137us bin size

## Slide 19: Production test with v1495 pulse generator
Jin-Yuan made a V1495 FPGA code which can drive V1495 as a pulse generator
The 40 MHz clock on the V1495 board is converted into 137 MHz (40*24/7) to drive all counters in the FPGA. (7.2917 ns)
The frequency does not match any frequency running on TDC card. So we can take it as random hit in some test.
We are using this pulse generator as a tool to perform the production test.
We prepared 4 kinds of test for Run2-TDC and two TDCs are tested.

## Slide 20: Test #1 Cross-Talk
Signal with time difference in the step of 7.2917 ns in each TDC channel.
TDC Channel (n) = TDC Channel (n+1) – 7.2917ns
TRIG

## Slide 21: Test #1 Cross-Talk

## Slide 22: Test#2 Resolution
If we fix the timing of the input signals and triggers, as we did in Test#1, all data are within 3 bins in the histogram. It is difficult to see how the calibration affects the TDC timing.
Example histogram from test#1

## Slide 23: Test #2 Resolution
A scan of hit timing in a step of 7.2917 ns.
For each event, trigger delay one clock cycle~ 7.2917 ns
Δt = t0 + n*7.2917 Mod(Δt,7.2917) = Mod(t0 + n*7.2917, 7.2917) = Mod(t0, 7.2917) = t0
TRIG
Mod

## Slide 24: Test2 without Bin Width Calibration

## Slide 25: Bin-Width Plot
There are 9 bins in 4 ns.
Each bin should be around 444 ps at ideal conditions.
We made calibration tables for trigger and all 64 channels.
ps
bin

## Slide 26: Test#2 with Bin-Width Calibration

## Slide 27: Test#2 Standard Deviation Study
Without Bin-Width Calibration, the σ is ~ 232.1ps
With Bin-Width Calibration, the σ is ~ 230.3 ps
Using channel 28 as an example
Without Calibration
With Calibration

## Slide 28: Test#2 with Bin-Width Calibration
Merging all the channels, the σ is 232 ps.
For the resolution that E906 needs, calibration won’t improve a lot. We may consider skipping the calibration.
Since we have the calibration table with both edge, we can always put calibration table into decoder if it’s required.

## Slide 29: Test#3 Multi-Hit
One trigger with four hits.
Each hit is 4*7.2917 ns after the previous one.
The level one buffer can record up to five hits in 64 ns.
5 hits per channel,
  Buffer pushing: clock cycle at 250 MHz.
Two Run2-TDCs passed this test
TRIG

## Slide 30: Test#4 Multi-Hit
One trigger with more than 8 hits.
Each hit is delayed by 72.917ns
Test level 2 & 3 Buffers
  Buffer popping: 1 hit/64 ns.
Two Run2-TDCs passed this test
TRIG

## Slide 31: Test#5 Multiple Hits Elimination
One trigger with four hits.
Each hit is ~29 ns after the previous one.
Set the Multiple Hits Elimination for 32 ns
Should get 2 hits per channel per event.
TRIG
Signal
Inhibit
TDC recorded

## Slide 32: Timeline
  One TDC is implemented in E906DAQ system and tested (run 2379).
  Feb 2013: new function development for multiple hits elimination and scalar
  Feb- Mar 2013: Production Test
    Test # 1~5 (Cross-Talk, Resolution, Multi-Hits for Buffer 1~3, multiple hits elimination)
    Full system - test VME noise feedback with 12 TDCs in one VME crate
  Mar 2013: full implement of TDCs with DAQ system and fine tune with E906 detectors/cosmic ray.

## Slide 33
That’s it~ thank you~

## Slide 35: Current Design
TDC internal clock
9 phase within 4ns
There is a free running clock inside the TDC (250MHz with 9 phases)
The TDC records the internal time for every hit in each channel and trigger in the first layer buffer. Up to 4 hits per channel can be temporarily stored.
After every 4 clock cycles at 62.5 MHz, the hits being read out are cleared. Data from a group of 4 channels are written into one of the 2 (or 4,8) circular buffers in the SRAM256x16 block waiting for a trigger.
When a trigger arrives, the writing address is pointed to the next circular buffer and the current circular buffer is read out.
TRIG
64 Channels Data

## Slide 36: 3 layers buffers
First layer buffer
second layer buffer
Third layer buffer

## Slide 37: First layer buffer
4 hits per channel,
  Buffer pushing: Any clock cycle at 250 MHz.
  Buffer popping: 1 hit/64 ns.
  RegPipe4a/ChGRP4a
Notes: nothing to do with trigger

## Slide 38: Second layer buffer
Second layer buffer, every 4 channels share a circular buffer
  (Page size)*(Number of Pages) = 256, user defined.
  Hits fill into a page continuously.
  When a trigger is received, hits are filled to next page and the current page is copied to the third layer buffer. The sequence is called copy in progress (CIP).
  (e.g. Page size=128 , CIP =128* 16 = 2048 clock cycle =2048 *16ns = 32us) (worth case scenario)
  SRAM256x16/ ChGRP16a
Notes: (CIP/ writing is trigger free/readout related to trigger)

## Slide 39: Second layer buffer
  SRAM256x16/ ChGRP16a

## Slide 40: Third layer buffer
Buffer size =1 K
  (Page size)*(Number of Pages) = 1024, user defined.
  Hit filling during CIP
  Readout from VME IRQ
  Outbuffer1k

## Slide 41
Test2 without Bin Width Calibration
Test2 with Bin Width Calibration

## Slide 43: Bin-Width Plot
ps
bin
ps
bin

## Slide 44
If Max hits per ROC is more than 434, the dead time will be large than 137us
In Run 2173 there are 2276/26225 = 8.7% events

## Slide 45
If Max hits per ROC is more than 391, the dead time will be large than 137us
In Run 2173 there are 4852/26225 = 18.5% events

## Slide 46: Circular Buffer Timing Window
The hits in the PipeReg4a are written to the circular buffer every 64 ns per hit.
If the PipeReg4a is filled with 4 hits, it takes 256 ns to write them into circular buffer.