## Slide 1: Silicon Photomultiplier Pixels in Vacuum and Cryogenic Conditions
Steve Kerby
3 August 2015

## Slide 2: Contents
Bulb Photomultipliers vs SiPM pixels
Silicon Photomultiplier Pixels: Behavior and Characteristics
Experimental and Analysis Setup
Initial Results
Dark Current, Signal Size, Wavelength Response v Temperature
Conclusions and next steps

## Slide 3: Bulb Photomultipliers
  Bulky
  Delicate
  1000+ V
  Damaged by large light leaks
Super-Kamiokande Accident (Nov 2001)
Hamamatsu

## Slide 4: Silicon Photomultiplier Pixels (SiPMs)
Compact
  3 x 3 mm
Resilient
  Survives light overflow
Low voltage
  ≈70 V operating level
Useful for arrays
Hamamatsu

## Slide 5: JEM-EUSO applications
Space-to-ground UHECP observation
Decreased background, increased efficiency, small pixel size
400 kHz dark current target

## Slide 6: SiPMs: Goals
Behavior in vacuum, at lower temperatures?
  Survival? Response?
  Dark Current Rate
  Signal Rate
  Wavelength Response
  Voltage Range Response
Usefulness of low temperature operation?
Useful for imaging array?

## Slide 7: SiPMs: Baseline Knowledge
Behavior at room temperature (≈20°C) well documented
Wavelength response provided by Hamamatsu, verified
Rated for use from 0° to 40°C

## Slide 8: Experimental Setup
Vacuum tube with SiPM array, 5x10-5 millibar, LN2 cooled
Light Source
Wavelength Selector
Sapphire Lens
Light Output Regulator

## Slide 9: Anatomy of a Scan: Analyzing the location and size of Photoelectron (PE) peaks
Logarithmic scale (ln(counts))
0PE
1PE
2PE
etc.
Gain

## Slide 10: Signal Rate Measurements
Rate: (Number of counts above threshold) * 333 Hz (Frequency of data collection software)
Signal Rate includes counts above 0.5PE

## Slide 11: Usable Voltages across SiPMs
Between Breakdown Voltage and Saturation Voltage
Lower range of voltages at lower temperatures
~4 V range at 20°C, 2 V range at -60°C
Chart: Signal Rate (kHZ) v Voltage, 20°C
| 69.043999999999997 | 68.646199999999993 | 68.248400000000004 | 67.8506 | 67.452799999999996 | 67.055000000000007 | 66.657200000000003 | 66.259399999999999 | 65.861599999999996 | 65.463800000000006 | 65.066000000000003 | 64.668199999999999 | 64.270399999999995 | 63.872599999999998 | 63.474800000000002 | 63.076999999999998 | 62.679200000000002
 | 9788.535 | 10343.313 | 11603.385 | 12881.772 | 15042.276 | 16506.144 | 14996.655 | 16319.664 | 15120.864 | 13132.521 | 10606.383 | 7328.997 | 2018.313 | 6.993 | 0 | 0 | 0
 | 8922.069 | 9166.491 | 10223.766 | 11392.596 | 13299.354 | 14608.044 | 13056.597 | 14217.768 | 13046.607 | 11215.107 | 8986.005 | 6255.072 | 1864.8 | 0 | 0 | 0 | 0
Chart: Signal Rate (kHz) v Voltage, -60°C
| 65.858599999999996 | 65.063000000000002 | 64.267399999999995 | 63.471800000000002 | 62.676200000000001 | 61.880600000000001 | 61.085000000000001 | 60.289400000000001 | 59.4938 | 58.6982
 | 12245.3424 | 11976.3117 | 13531.6548 | 16156.2276 | 21819.2589 | 17265.5172 | 12460.0275 | 4239.2565 | 3.996 | 2.1978
 | 12342.1455 | 12049.6383 | 13810.176 | 16522.1613 | 20220.1596 | 15227.0577 | 10673.7156 | 2212.8849 | 39.96 | 59.94
Different signal rates due to improved light source setup at low temperatures, shift in range of voltages independent of light source setup

## Slide 12: Room Temperature Behavior, 22°C
Vacuum conditions have no effects (fortunately)
Dark Count Rate ~ 1600 kHz
1+ PE peaks despite no light source
Noticeable 1+ PE rate with LED (1600 kHz => 2800 kHz)
Light Source On
Light Source Off
Logarithmic scale

## Slide 13: Cooled Behavior, -20°C
Dark Count Rate ~ 35 kHz
Visible addition of 1, 2, 3 PE peaks
Significant rate difference, 35 kHz => 7000 kHz*
*Different light source setup from room temperature measurements leads to higher signal rate
Light Source On
Light Source Off
Logarithmic scale
Note: Different ADC scales

## Slide 14: Cooled Behavior, -60°C
Dark Count/Signal Ratio more pronounced
Dark Count Rate ~ 6 kHz
6 kHz => 6000 kHz
Light Source On
Light Source Off
Logarithmic scale
Note: Different ADC scales

## Slide 15: Dark Current vs. Temperature
20°C
-20°C
-60°C
Significant decreases in dark current
1600 kHz => 35 kHz => 7 kHz

## Slide 16: Dark Counts vs. Temperature, Cont.
0.6V above breakdown voltage for each temperature
<400 kHz dark counts goal surpassed at 10°C
Slightly cooling SiPM pixels decreases Dark Counts significantly
Chart: Dark Current Rate (kHz) v Temperature (°C)
| 20 | 15 | 10 | 5 | 0 | -5 | -10 | -15 | -20 | -25 | -30 | -35 | -40 | -45 | -50 | -55 | -60 | -65 | -70 | -75 | -80
 | 791.208 | 459.54 | 321.012 | 247.752 | 128.205 | 91.908 | 73.926 | 56.277 | 23.976 | 22.311 | 13.653 | 17.316 | 12.987 | 11.322 | 10.656 | 7.659 | 8.325 | 9.324 | 7.659 | 6.66 | 5.661
 | 770.895 | 450.549 | 356.643 | 254.745 | 140.526 | 150.183 | 82.584 | 41.958 | 36.297 | 13.986 | 15.651 | 13.653 | 10.656 | 14.319 | 8.991 | 9.324 | 4.662 | 6.993 | 6.993 | 6.327 | 7.326
 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400 | 400
400 kHz JEM-EUSO spec

## Slide 17: Signal Rate vs. Temperature
Signal rate constant with decreasing temperature
Increased Signal/Noise ratio at lower temperatures
Decreasing temperature does not decrease signal receptivity
Chart: Signal Rate (kHz) v Temperature
| 22 | 15 | 10 | 5 | 0 | -5 | -10 | -15
 | 14719.7655 | 17211.6045 | 12962.1915 | 18105.0435 | 13093.7265 | 16728.7545 | 16638.345 | 16741.7415
16000 kHz (In these conditions)

## Slide 18: Wavelength Response v. Temperature
At -30°C, response spectrum resembles room temperature spectrum
Decreasing temperature does not change wavelength response spectrum
Chart: Relative Response v Wavelength (nm), -30°C
| 380 | 480 | 600 | 700 | 800 | 900
 | 0.58749 | 1 | 0.76102 | 0.34106 | 0.39146 | 0.28199

## Slide 19: Conclusions
Room temperature behavior matches reported specs
Cooled behavior
  Significantly less dark current
  Consistent peak wavelength 400-500 nm
  Predictable, constant signal rate
  Extremely low temperatures not necessary

## Slide 20: Conclusions, Cont.
  Relative ease of reducing dark current rate
  Consistency of wavelength response, signal rate
There are immediate benefits to SiPM operation at low temperatures
These benefits are accessible at only slightly cooled temperatures

## Slide 21: Next Steps: Imaging Array
Reducing dark current with cooling
Checking cross talk rates
Testing with precision light image sources
Precision software triggers and calculations
SenseL

## Slide 22: Thanks for Listening!
Special Thanks to Juan Estrada, Kevin Kuk, Lee Scott, Paul Rubinov
sekerby@fnal.gov or sek93@case.edu
Questions?