## Slide 1: Status of St-4 Muon Prop Tubes 05/27/2010 E906 Collaboration Meeting
Ming X. Liu
LANL
Gas test
Module test
Installation and commissioning plan

## Slide 2: Progress at Fermilab
03/4-7: fixed gas tubing, HV, signal lines …
  Imran, Andrew, Ming…
04/19-24: test gas & fix broken modules
  Lei, Chuck, Ming et al.
05/9-16: HV scan, first time 8-ch readout
  Imran, Yen-Chu, Kaz, Chuck, Ming et al.
4/10
5/10
3/10
3/10
3/10

## Slide 3: Chicago’s Famous Pizza … and Spaghetti

## Slide 4: Gas Test - Summary
Maximum drift time study:
  MIPP: 0.35uS
  Ar/CO2 (80/20): 2.5uS
  Ar/CO2(95/5):1.5uS
  D0 Muon Gas: 0.8uS
HV plateau scan
  MIPP: 1850-2100V
  Ar/CO2(80/20): 1900-2100
  D0 muon gas:1900-2100
D0 Gas
Ar/CO2 80/20
MIPP Gas

## Slide 5: MIPP Gas: It is Fast! ΔTdrift ~ 0.35 uS @2000V
Module

## Slide 6: Maximum Drift Time Measurements Module #34, chan-9..
ΔT=0.35 uS
Trig.
Signal
Coin. Gate
Time

## Slide 7: MIPP Gas

## Slide 8: Qualitative HV scan and Signal Rates with Cosmic Rays (Ar/CO2)
Cosmic trigger

## Slide 9
Bad Threshold Resistor
Module 14 Module 23 Module 20 Module 43 Module 35 Module 100
Module 2 Module 36 Module 38 Module 40 Module 103 Module 17
Module 15 Module ? Module 9 Module 7 Module 37 Module ?
Module 33 Module 30 Module 105 Module 45 Module 51 Module 1
Module 11 Module 26 Module 29 Module 28 Module 31 Module 32
Module 10 Module 50 Module 24 Module 42 Module 27 Module 25
Module 46 Module 16 Module 104 Module 44 Module 101 Module 4
Module 47 Module 52 Module 102 Module 48 Module 18 Module 6
Good, Vmax>2050V
Dead
Noisy/Background/high Vmin//low Vmax
Unusable module
Module 34
Module 41
No PreAm, Flowing gas now 4/23/10
16 15 14 13 12 11 10 9
8 7 6 5 4 3 2 1
NIM Channel map

## Slide 10: Quantitative HV Plateau Scan with 8-ch Readout for all Tubes (May 2010)
Tested all individual tubes w/ DAQ, 8-channel scalar readout (Thanks Yen-Chu, Kaz and Josh!)
  Imran, Lei, Chuck, David C., David N., Ming …
  Ar/CO2 for most of the modules
  One stack with MIPP fast gas (Mod. #34)
    CH4/CF4/Ar = 8.5% / 15% / bal.
  One module with D0 muon gas (Mod. #17)
    CO2/CF4/Ar = 8% / 8% / bal.

## Slide 11: HV Scan with 8-ch Readout ECL to ECL scalar

## Slide 12: DAQ 8-ch Scalar Readout Screen Dump

## Slide 13: HV scan: MIPP Gas for Module #34
1850-2150

## Slide 14: HV scan for Mod. #38 Ar/CO2 (80/20)
1850-2100
Broken tube

## Slide 15: Good Modules from Quantitative HV scan
 |  | 41 |  | 34 | 
14 | 23 | 20 | 43 | 35 | 100
2 | 36 | 38 | 40 | 103 | 17
15 | ?? | 9 | 7 | 37 | ??
33 | 30 | 105 | 45 | 51 | 1
11 | 26 | 29 | 28 | 31 | 32
10 | 50 | 24 | 42 | 27 | 25
46 | 16 | 104 | 44 | 101 | 4
47 | 52 | 102 | 48 | 18 | 6
Good modules (mostly)
Draw high current above 2000 V due to a couple of bad channels.
Still need to be tested. (different shape modules).
Draw very high current.
Some unstable channels.

## Slide 16: Number of Good Modules
Need 2x8+2x9 = 34 modules
Green ones: 24+7 =31  we need 3 more good ones!
Possibly good ones:
  4 Modules:
    # 6, 32, 42, 101 were tested good in April 2010. High current problem probably due to poor gas and/or humidity, as we stopped flowing Ar/CO2 gas for a while for these modules.
  2 modules:
    # 18 and #43, need to replace preamp board (#43) and preamp card(#18).
  Module #100
    flaky FEM board and poor grounding?
  7 wrong shaped ones
    Can we use 4 of them, but read them out from the opposite side?
Would like to run them all with MIPP gas and check their performances
We will have enough good modules if we can recover these possibly good modules – very likely.

## Slide 17: Picture Frames and Installation Plan
Initial design done, under Engineering evaluation at LANL
  Thanks Tom, David N., Walt/John/Matt (LANL)
Got a Quote from 80-20 for the materials based on initial design
  Thanks Tom!
Once Paul and Chuck obtain the green light from Fermilab, I would like to place the order for materials ASAP (some time in June)
Assembly and installation in late June - July

## Slide 18: The Design of the Picture Frames (2nd version)

## Slide 19: More Details about Picture Frame

## Slide 20: Quote from 80-20
4x1.8K =$7.2K
LANL Engineers’ efforts $xx

## Slide 21: Future Plan
Recover “bad” modules:
  Test 7 wrong shaped modules with Ar/CO2 at next visit, try to use 4 of them if possible.
  Test 4 modules that were tested good but failed at last checkup - likely due to bad gas.
Would like to run “good fast gas” for all modules, scan for HV plateau for all tubes, identify 34+ good modules for the final assembly
  MIPP or D0 muon gas (or something similar)
Assembly and installation:
  Get Fermilab structural engineers’ approval ASAP, order the materials from 80-20 in June.
  Assembly and installation in late June/July
Electrical ground and noise
  Need to run the full DAQ and some sort of online/offline analysis code
  Need ECL to ECL TDC, 34x16=544 channels
Would like to get the detector in for the test beam and check the background rates this summer