## Slide 1: IB stave construction and prototypes
ALICE ITS Upgrade
Stave Production Readiness Review
CERN, 27 April 2017

## Slide 2: Introduction
The stave assembly procedure has been established with the aim to ensure:
  chip integrity (IB silicon chips are thinned to 50 mm)
  precise alignment to better than 10 mm
  reliable and reproducible results
The main steps are:
  Chips alignment
  Chips gluing on FPC (HIC) and wire bonding
  HIC gluing on carbon fiber spaceframe (STAVE)
After the EDR, we focused mainly on 2) since 1) is ensured by the MAM and 3) was already well advanced

## Slide 3: Interconnection layout
silicon
glue
Upilex-75S
FPC
coverlay 20 um
Al 25 um
p.i. 75 um
Al 25 um
coverlay 20 um
Ablestik 45W 90 um
ALPIDE 50 um
Hybrid Integrated Circuit (HIC) schematic cross section
Main constraints for FPC to silicon gluing:
Optimize/minimize amount of glue
Ensure good adherence under FPC pads, no vibrations during wire bonding
Prevent glue contamination of pads to be bonded
Optimisation of gluing mask
Characterisation of glue layer
(power supply connection)

## Slide 4: Jigs for HIC assembly
FPC gripper
Recess to accommodate SMD and wire bonds
Planarity: 15 um
FPC extension support
(temporary)

## Slide 5: Jigs for HIC assembly
Chips vacuum chuck
(HIC assembly table)
Planarity: 13 um
Wire bonding table

## Slide 6: Gluing mask optimisation
Initially (at the EDR) the procedure to glue FPC and chips was based on adhesive mask Oracal 638, 90 um thick associated to shimming producing a final glue layer of 50 um, compatible with the sum of tolerances involved
Observed issues:
  Glue spots not close enough to vias edges
  Weak adherence
Coverlay
Silicon chip
Polyimide (Upilex 75S)
Aluminiium
Aluminium
Coverlay
Ablestik 45
Chip vacuum chuck
FPC gripper
Assembly stack-up
Component | Tolerance
Chip vacuum chuck planarity | ± 7 um
Chip thickness | ± 5 um
FPC thickness | ± 10 um
FPC vacuum chuck | ± 8 um
TOTAL (linear) | ± 30 um
TOTAL (quadratic sum) | ± 15.5 um
Glass plate glued on chip using mask with 50 um shimming

## Slide 7: Gluing mask optimisation
Upgraded gluing procedure:
  Optimised holes amount and geometry (remove holes where not necessary, use everywhere Ø 0.7 mm)
  SDAG adhesive film RS-150-137 (150 um thick) with 90 shimming (Oracal 638 tape)
  Hand mixing (3g for 30”) instead of mixer dispenser nozzle

## Slide 8: Glue characterisation
Chemical analysis for ionic impurities (MIL STD 883H): dangerous ions concentration is below threshold
Veritas laboratories (F)
Pull tests
Glue prepared by hand mixing for 30”
5 samples of FPC + single chip using last procedure (mask/shimming)
Average load measured on 3 samples : 3.9 kg
2 samples under ageing test
ADD RESULTS LATER
+ TEST ON PIQ CHIPS
ADD PHOTO OF SAMPLE

## Slide 9: HIC assembly procedure overview
Chip automated inspection in ALICIA
Chip alignment on ALICIA assembly table
FPC clamping by vacuum on dedicated jig, alignment using dowel pins
Adhesive mask application on FPC and glue distribution
Mask removal, check of glue spots, application of shims (adhesive tape) on FPC wings
Flip FPC and overlap on chips using ruby balls for alignment
Curing on assembly table at room T for 12 h
Transfer on wire bonding jig
Visual inspection
Chips interconnection by wire-bonding
Transfer on HIC storage plate and electrical test

## Slide 10: 1. Chips inspection
Chips are identified by the serial no. of the storage tray and by the position inside it. Each position in a tray is uniquely associated to a position in the original wafer
Prior to assembly, chips are inspected to verify dimensions (the size shall not differ from nominal one by more than 5 mm), edge integrity (no cracks shall be present at a distance shorter than 10 mm from the seal ring) and pad cleanliness (the total area occupied by one or more particles shall not exceed 10% of the pad surface).

## Slide 11: 2. Chip alignment on assembly table vacuum chuck
Chips which passes the quality inspection are aligned on the vacuum chuck of the assembly table
An accuracy of ± 5 mm is required on the position; the gap between the chips is 150 mm (seal-ring to seal-ring)
After alignment a further visual inspection is performed to spot presence of bumps produced by particle under the chips (→ chips and table cleaning)

## Slide 12: 3. FPC mounting on gripper
The FPC is aligned on the gripper using dowel pins and vacuum is started

## Slide 13: 4. Adhesive mask application on FPC and glue distribution
Glue preparation
The gluing mask is aligned on top of the FPC using the same dowel pins
Glue distribution using a metallic spatula
1
2
3

## Slide 14: 5. Mask removal and application of shims (90 um tape)
1
2
3

## Slide 15: 6. Flip FPC and overlap on aligned chips using ruby balls
1
2
3
4
5

## Slide 16: Last steps
7. Curing at room T for 12 h
8. Removal from machine and transfer on wire bonding jig (copy of chips vacuum chuck)
2
3
4
5
1
6
9. Visual inspection (alignment, presence of glue)
10. Chips interconnection by wire-bonding
2
3
4
1

## Slide 17: 11. Transfer on HIC storage plate and electrical test
2
3
4
1
New base plate of bonding jig with FPC extension support
New FPC extension support on FPC gripper
Spring loaded pin to align/release FPC/HIC
HIC storage plate
New tools under procurement

## Slide 18: Summary of IB HIC prototypes
HIC ID | Gluing procedure | Chip | status | Comments
IB-HIC-proto-WB05 | 50 um shims, mask V4 | 100 um | working | 3 bonding wires, no shorts
IB-HIC-proto-WB08 | = | = | working | 3 b. w., initial shorts
IB-HIC-proto-WB09 | = | = | working -> STAVE-0 (NA61) | 1 b. w., initial shorts
IB-HIC-proto-WB10 | + hand mixing | = | partially working | 1 b. w., initial shorts
IB-HIC-proto-WB11 | = | = | working | 1 b. w., initial shorts
IB-HIC-proto-WB12 | = | = | partially working | 1 b. w., initial shorts
IB-HIC-proto-WB13 | + mask V5 | = | working | 1 b. w., initial shorts
IB-HIC-proto-WB14 | 90 um shims + mask V6 | = | working, ageing test done | 1 b. w., no shorts
IB-HIC-proto-WB15 | + mask V7 | = | working, ageing test done | 1 b. w., no shorts
IB-HIC-proto-WB16 | = | = | working, ageing test done | 3 b. w., no shorts
IB-HIC-proto-WB17 | = | 50 um | working, ageing test partial | 3 b. w., “small” shorts
IB-HIC-proto-WB18 | = | = | working -> STAVE-1 | 3 b. w., no shorts
IB-HIC-proto-WB19 | = | = | working, ageing test in progress | 3 b. w., no shorts
IB-HIC-proto-WB20 | = | = | 8 chips working | 3 b. w., no shorts, chip 9 broke after transfer on test plate
List of prototypes produced from 08/2016

## Slide 19: Summary of IB HIC prototypes
HIC ID | Gluing procedure | Chip | status | Comments
IB-HIC-proto-WB21 | 90 um shims + mask V7 | 50 um | working -> STAVE-2, ageing test in progress | 3 b. w., no shorts
IB-HIC-proto-WB22 | = | 100 um PIQ | working, ageing test in progress | 3 b. w., no shorts
IB-HIC-proto-WB23 | 50 um shims + mask v7 | 100 um PIQ | working | 3 b. w., initial short in one chip
IB-HIC-proto-WB24 | 90 um shims + mask V7 | 50 um | working | 3 b. w., initial shorts
IB-HIC-proto-WB25 | = | 50 um | working | 3 b. w., initial shorts
IB-HIC-proto-WB26 | 50 um shims + mask v7 | 100 um PIQ | under test | 
IB-HIC-proto-WB27 | = | 100 um PIQ | under test | 
List of prototypes produced from 08/2016

## Slide 20: Wire bonding characterisation
Results of pull test on 80 bonds
  Min load= 8.278 g, Max load= 13.476 g, avg = 11.005 g, stdev = 1.08 g
  Avg- 3stdev = 7.764 g (> 7 g min. ref. load)
  Grade details:
    1st bond heel break , 79 events, avg. 10.993 g
    2nd bond lift, 1 event, 11.926 g
RESULTS FROM AGEING TEST TO BE ADDED

## Slide 21: STAVE assembly procedure

## Slide 22
IB STAVE ASSEMBLY
ASSEMBLY PROCEDURE
TOOLING AND MECHANICAL TOLERANCES
To be updated

## Slide 23
Glue (50µm)
Pixel Chip (50µm)
5 mm
Cold plate
FPC (175µm)
15,4 mm
INNER BARREL
Layout
16.15 mm
15 mm
0.3 mm
0.85 mm
To be updated

## Slide 24
IB SF&CP + HIC Gluing
Objective:
glue the HIC to the IB spaceframe & coldplate, such to provide HIC support, alignment, and thermal contact to the coldplate
temporary test wires
To be updated

## Slide 25
Assembly components and jigs
jig used for HIC gluing to IB spaceframe
SF_ jig
HIC_ jig (vacuum)
IB Spaceframe& Coldplate
glue
FPC
To be updated

## Slide 26
IB SF&CP + HIC Gluing steps
1. Install the IB speceframe&coldplate in the SF_ jig
To be updated

## Slide 27
IB SF&CP + HIC Gluing steps
2. Position gluing mask (100 µm thick), use locating pins
3. Apply glue ECCOBOND 45 and level with spatula
4. Remove gluing mask
To be updated

## Slide 28
10 µm
25 µm
75 µm
25 µm
10 µm
100 µm
500 µm
IB SF&CP + HIC Gluing steps
HIC_jig (vacuum chuck)
HIC_jig has to provide room to allocate the FPC capacitors and bonded wires
5. Position the HIC on the HIC-Jig, use locating pins and vacuum
To be updated

## Slide 29
6. Position shims to accommodate for a nominal 50 µm glue thickness after the matching of the 2 jigs
IB SF&CP + HIC Gluing steps
To be updated

## Slide 30
IB SF&CP + HIC Gluing steps
7. Position the SF_ jig on top HIC_jig, use locating pin and fasteners.
8. Rotate the assembly and cure 24h at RT.
To be updated

## Slide 31
IB SF&CP + HIC Gluing steps
9. Remove SF_jig
10. Remove Stave from SF jig and HIC_jig
To be updated

## Slide 32
11. Remove wings – cuts wing bridges, procedure under test
IB SF&CP + HIC Gluing steps
To be updated

## Slide 33
IB SF&CP + HIC Gluing steps
12. Acceptance Tests
Test setup for first prototype
To be updated

## Slide 34
IB STAVE ASSEMBLY
ASSEMBLY PROCEDURE
TOOLING AND MECHANICAL TOLERANCES
To be updated

## Slide 35
Tooling and mechanical tolerances
Requirements:
Geometrical tolerance 10 micron (planarity)
Measured:
Geometrical tolerance 14 micron (planarity)
~12µm
~14 µm
To be updated

## Slide 36
32 µm
~21µm
HIC
Spaceframe+
Connectors
Old jigs, out of tolerance
Tooling and mechanical tolerances:
engineering jig (old)
To be updated

## Slide 37
Spaceframe 1_21 ( w/o connectors) free
With HIC, stave 1_21 free
With HIC, stave 1_21 clamped
+0.033 mm
- 0.033 mm
+0.034 mm
- 0.034 mm
+0.039 mm
- 0.039 mm
free
free
clamped
clamped
free
free
screw
ruby
Before HIC gluing
After HIC gluing
Tooling and mechanical tolerances:
tol achieved with engineering jig (old)
To be updated

## Slide 38: Production workflow

## Slide 39: Production planning