## Slide 1
sPHENIX LDRD Review
UNCLASSIFIED
Xuan Li
P-25
ALPIDE Introduction and Test Stand
Outline
ALPIDE chip introduction
LDRD test bench setup
LDRD test bench achievements
Summary and Plans

## Slide 2
The MVTX Basic Unit: ALPIDE sensor
ALPIDE chip Introduction:
The 3-layer MVTX consists of 48 staves and each stave contains 9 ALPIDE chips.
Stave
Scheme of ALPIDE (524288 channels)
3cm
1.5cm

## Slide 3: MVTX Performance Requirements
w/ perfect acceptance
no pile up
w/ perfect acceptance
no pile up
Can MVTX meet these requirements in real data?
  Small dead channel fraction High efficiency.
  High signal/noise ratio  Low background.
The heavy flavor program at sPHENIX led by the MVTX requires good vertex resolution and high efficiency.

## Slide 4
LDRD Test Bench Components
Power Board
Filtering Capacitors
Samtec Twinax “FireFly”
9 Data (1.2Gbps)
1 Clock, 1 Control/Trigger
Readout Unit
Regulated Power
FELIX
Data (9.6 Gb/s max)
Control
Trigger
Alpide Sensors
FPC
Cold Plate
regulators
Power
5 m
Optical Links
9 Sensor Stave
RCDAQ
Purchased one stave and 5 individual ALPIDE chips.
Test bench
readout module:
MOSAIC
Purchased 1 Power Board

## Slide 5: ALPIDE Characterization: Good Active Area
Measure the dead areas for efficiency evaluation.
Test with internal pulse generator:
  Scanned available chips and stave at LANL through digital scan to verify the good channel fraction: the active pixel fraction is >>99%.
  Similar results with different readout speed.
No. of 0 hits
777
Cold channel
1
Hot channel
0
Bad fraction
0.15%
ALPIDE chip test example

## Slide 6: ALPIDE Characterization: Low Noise (I)
Noise scan per chip
Signal/Noise ratio scan
Charge injection with internal pulse generator:
  Scanned available chips and stave at LANL through threshold scan: signal/noise based on charge injection study is > 60. Comparable with CERN test (~60).
Proves the background contribution is low!
Chip test with internal trigger
Signal scan per chip
Noise scan per chip
Signal/Noise uniformity
MIP~1000 e-
Signal/Noise scan per chip

## Slide 7: ALPIDE Characterization: Low Noise (II)
Charge injection with internal pulse generator:
  Scanned available chips and stave at LANL through threshold scan: signal/noise based on charge injection study is around 60. Comparable with CERN test (~60).
Stave test with internal trigger:
Simultaneous readout of all chips.
Noise
Signal/Noise
Inactive chips due to
communication failure.

## Slide 8: ALIPIDE Performance Evaluation Configuration
Study the ALPIDE performance with real particles.
Tests in triggered mode with default delay at 2μs:
  External trigger with cosmic ray muons (rate: ~0.1HZ).
  External trigger with 90Sr beta decay electrons (rate: ~1Hz).
Scintillator 1
Scintillator 2
ALPIDE chip
e-
Sr90
Scintillator 1
Scintillator 2
ALPIDE chip
μ-
90Sr setup
Electron maximum
energy 2MeV
Cosmic ray setup
Muon average
energy 2GeV

## Slide 9: Test Achievements: verify trigger timing and signal detection
Time in trigger and observe physical signal.
For example: 90Sr source test.
Collimator
ALPIDE chip
Row 0
Row 511
Collimator
90Sr source
ALPIDE
chip
Scintillators
Pixel hit 2D map
Profile along column
Profile along Row
Clear collimator structure seen in the ALPIDE chip.
Row
Column
Column
Row

## Slide 10: Test Achievements: good hit resolution
Scintillator 1
Scintillator 2
MAPS chip
μ-
Initial look at tracking with two layers of ALPIDE chips.
Found clusters that belong to a cosmic ray muon track.
Test stands
Event Display
Characterize the cluster size: related to the hit resolution.
Cosmic ray
Linear scale
Comparable with 90Sr source test results and DESY beam test (cluster size < 3 pixels).
Average cluster size (2 pixels) indicate the hit resolution < 10μm  Meets requirements.

## Slide 11: Under study: Trigger delay latency
Can the MVTX adapt to the sPHENIX trigger scheme with default trigger delay at 4-5 μs?
The default ALPIDE trigger delay is 2μs.
Need to tune the pulse shaping and trigger delay.
Analog IN
After lowering the threshold to optimize the pulse shaping time:
No significant performance differences with trigger delay from 2μs to 6μs.
Digital IN
Trigger
Trig Delay
Duration:5-10μs

## Slide 12
Integration: Power Board Test
Power Board
Filtering Capacitors
Samtec Twinax “FireFly”
9 Data (1.2Gbps)
1 Clock, 1 Control/Trigger
Readout Unit
Regulated Power
FELIX
Data (9.6 Gb/s max)
Control
Trigger
Alpide Sensors
FPC
Cold Plate
regulators
Power
5 m
Optical Links
9 Sensor Stave
RCDAQ
Successfully operating the ALPIDE using the Power Board prototype!
MOSAIC GUI for PB

## Slide 13: Summary
Initial results achieved from the LDRD test bench:
  >>99% good channels on chip  High efficiency.
  MIP/Noise ≥ 200  Low background.
  Hit resolution from cosmic ray or 90Sr beta decayed electrons < 10μm  Meets the MVTX requirements.
  Proof of principle study proves the feasibility to adapt the MVTX to the sPHENIX trigger scheme.
  Successful test of the Power Board prototype as part of the MVTX electronics.

## Slide 14: Plans
Set up the laser test bench to study the efficiency, noise rate, trigger delay time, threshold, pulse shaping, etc.
X-Y scan table
Laser generator:
Duty cycle: < MHz
Focus area: < 10μm2
μ-
Telescope for cosmic/source test
scintillator
ALPIDE
chips
scintillator
Set up the telescope containing multiple layers of ALPIDE chips (staves) to study the track resolution, efficiency, etc, from cosmic ray and source tests.
Complete optimizing the ALPIDE performance with the RU+FELIX+PB full readout chain this summer.

## Slide 15: Backup
ALPIDE pixel block diagram.

## Slide 16: Backup
Cluster evaluation by the CERN team from 450 MeV e- beam test.
Cluster Size
Cluster Profile along X
Cluster Profile along Y
Even particle energies are different, similar results are achieved between the test bench studies with cosmic ray or 90Sr beta decayed electrons.

## Slide 17: Achievements: cluster determination
Cluster size determined from 90Sr beta decayed electrons.
The average cluster size determined from 90Sr source is comparable with cosmic ray test.