## Slide 1: MVTX Commissioning Plan
09/15/2021
Ming Liu for the MVTX Team
LANL
MVTX Commissioning Plan Live Google Doc:
https://docs.google.com/presentation/d/1CYn0g2FkmQzHUbXfIOD8elDWMbOzizJn8zK36RyVFvA/edit?usp=sharing

## Slide 2: MVTX Detector: inner-most sPHENIX subsystem
Service Barrel (SB)
Cylindrical Shell
Structure (CYSS)
End-Wheels (EW);
- 3 layers
3-layer sensor barrel
- 48 staves, 432 chips
R (mm) | min | mid | max
Layer 0 | 24.61 | 25.23 | 27.93
Layer 1 | 31.98 | 33.35 | 36.25
Layer 2 | 39.93 | 41.48 | 44.26
North
North
New BP
Cantilevered detector: only on the S-side of sPHENIX
Carb. Fib. Comp./CFC = EW, CYSS, SB
Total Weight: ~24 kg (sensors: ~3 kg)
Sensor Length: 271mm
Notes: A quick reminder of the MVTX overall design -

## Slide 3: MVTX Milestones
Carbon Fiber Components: completed 7/2021 (Workshape/France)
  Assembly fixtures: designed
  Assembly manual: completed
Stave production: completed 5/2021 (CERN)
Half-detector assembly: ongoing (LBNL)
  CFC CMM & test fitting: since July, 2021
  Half-detector assembly review: 09/09-10, 2021
  Complete assembly: by Feb. 2022
Aim to meet the original MVTX installation dates
- MVTX pre-installation commissioning at BNL: March- Sept., 2022
- IR available for MVTX installation: Sept. 23, 2022
Pre-COVID19 Schedule in MVTX PMP
Assembly review, 9/9/2021
Notes: Following the approval from this review, we will start CFC production from Jan 2021 …. We have 40+ gold staves at CERN waiting for shipping .. Need 48 for the final detector

## Slide 4: MVTX IR Installation Schedule (sPHENIX Plan)
MVTX:
9/23 – 10/18,
2022
Complete pre-installation commissioning at BNL before 9/23/2022

## Slide 5: MVTX Commissioning Plan @BNL
Setup clean tent for half-detector QA
  Jan. – Feb.,2022
Pre-installation commissioning
  Mar. – Sept. 2022, ~ 6 months
IR installation
  9/23-10/18, 2022, ~3 wks
Post installation commissioning
  10/19 – 12/31, 2022, ~8 weeks

## Slide 6: Pre-installation @ 1008: Jan. – Sept., 2022
Setup clean tent at 1008
  Laser room for detector and assembly work?
  other services stay outside of clean tent
MOSAIC test stand for stave QA
  PU+MOSAIC reside in a VME Crate
  2 test stands
QA two half-detectors
  Single stave QA, one-by-one, 48 total
  48 SamTec readout cables
  48 DVDD cables
  48 AVDD cables
  48 “HV” bias cables
  48 PT100 T-readback
QA time: 4hrs/stave
  Two MOSAIC systems available, ~24 days
MOSAIC
Test Stand
MVTX Detector

## Slide 7: Single Stave QA
MOSAIC test stand
  MOSAIC board with custom SamTec readout cables
  Power Unit for stave with custom power cables
  Bench PS for PU
    +3.3V, -5V(BB)
Cooling system
  Negative-pressure (about -30%)
  Ambient temperature
QA:
  Dead/hot pixels
  Threshold scan, noise level
  Error rates
chiller
stave
MOSAIC
PU,PS

## Slide 8: Possible re-work on half-detector cables and staves

## Slide 9: Stave Assembly Fixtures: to replace bad staves
Notes: Special assembly jigs are required to put staves on the EW on exact locations to provide prevision tracking measuremetns

## Slide 10: Assembly Area Setup at LBNL 1)Assembly jigs; 2)CMM inside a clean tent ~ similar setup needed at BNL
CMM
Assembly jig
Clean tent

## Slide 11: Detector Assembly Procedures … CMM after each step (~10 um tolerance)

## Slide 12: Pre-Installation Stretched Goals
Small scale full MVTX readout chain (if not the full half-detector)
  A minimal “full readout chain”: (MVTX = 6 x )
    8-stave + 8-RU + 4-PU + 1-FELIX + 1-mGTM + 1-AMD
    Use the final long (~12m) readout and power cables (from PP3 to Rack 2E1)
    Final CAEN power system (requires 208V, 3~)
    Final cooling system
      Stave cooling
      RU/PU cooling
    System endurance test
QA software and online monitoring
    Modify ITS QA software for MVTX
Cosmic ray runs
  Channel mapping
  MVTX internal alignment

## Slide 13: MVTX IR Installation: 9/23 – 10/18, 2022
MVTX X-Wing
Pre-installation dry run at BNL (MIT)
full mockup @BNL
beam pipe, MVTX, INTT/TPC volumes …
Dates: early 2022
Notes: And the detector is supported through X-wing that is attached to iHCal supporting structure s..

## Slide 14: MVTX Detector Installation and Interface
sPHENIX-SE-ICD-042 Document:
https://docs.google.com/spreadsheets/d/1hHUhxZOZv8WEs3FH361_gN6McVU4Zhy60sGtYotRHMs/edit?usp=sharing
Power and safety interlock
PP3: all service lines
PLC Remote IO

## Slide 15: Post-Installation Commissioning
Full MVTX system readout, slow control and QA
  Calibration, threshold scan etc.
  Online monitoring and control, WinCC-based
Adapt many of ITS-2/IB commissioning tools
  WinCC platform
  Modify ALICE QC software for sPHENIX
Detector safety interlock
  V, I, T, P etc.
Integrate MVTX into the sPHENIX DAQ
  Readout with the rest of sPHENIX subsystems
  Triggered and CR modes
  Offline event reconstruction (INTT+TPC+…)
Cosmic ray runs
  Continuous Readout
    MVTX internal alignment
  External triggers (TPOT, EMCal/Hcal possible?)
    Global alignment

## Slide 16: MVTX ~ ALICE ITS Surface Commissioning - detector control and status monitor
Control and Monitoring Plots from ALICE/ITS-2 commissioning runs:
WinCC – based system

## Slide 17: MVTX similar to ITS Surface Commissioning - check each chip’s hit/noise occupancy, errors
Layer-0: 12 staves
Layer-1: 16 staves
Layer-2: 20 staves
staveID
chipID
chipID
chipID
staveID
staveID

## Slide 18: ITS Surface Commissioning: Check out all chips - threshold, dead/hot pixels in each chip
# of e-
Dead pixles

## Slide 19: MVTX Internal Alignment: Detector QA Data
Work in progress:
update MVTX code to read in new GDML file for each stave placement
allows stave-specific alignment, such as Chip placement (~+/-10um tolerance@assembly), from stave production QA DB
Create GDML file for each stave:
ML based detector alignment:
MC samples generated with misaligned staves
ML package used to find correction factors

## Slide 20: Manpower at BNL
LBNL
LANL (plan to have one PD station at BNL)
ORNL
MIT (“local”)
BNL (“local”)
other institutions
extended travels to BNL by experts:
students/postdocs/scientists, engineers, technicians…
Try to maintain at BNL at least:
~ 2 FTE + 2 PD + students
from MVTX project (2017) file for “installation”:

## Slide 21: Additional slides

## Slide 22: Assembly work at LBNL (may need to do similar work at smaller scale)

## Slide 23: Half Detectors

## Slide 24: Commissioning Plan @BNL in 2022 (Draft)
Pre-installation @1008: Jan. – Sept., 2022
  Setup clean tent @1008, Jan-Feb, 2022
  Check out half detectors inside clean tent
    Single stave QA with MOSAIC test stand
    Fix issues
  Full readout with RU+FELIX
    Minimal 8-stave + 8 RU + 1 FELIX + AMD Server
    Cooling system
    Stress and endurance test
  Slow control and monitoring
  CAEN power system test
  Full cooling system test
  Cosmic runs for channel mapping and alignment etc.
MVTX installation at IR: 09/23-10/18, 2022 (3 wks)
  Insertion of MVTX in IR
  Readout system and cabling
  CAEN power system and cabling
  Install cooling system
  Detector safety system installation
  System test
Post-installation: 10/18 - 12/31, 2022 (8 wks)
  Full system test and debugging
    Readout and slow control
    Cooling
    Detector safety interlock
  Cosmic ray runs
  Online monitoring and controls
  MVTX stave internal alignment
  MVTX Global alignment

## Slide 25: MVTX Milestones and Key Tasks (from MVTX PMP 2019)
From MVTX PMP
12/2019
Review today:
09/09/2021
We aim to finish all
ahead of the schedule

## Slide 26: MVTX Half-Detector Test @LBNL
Test staves at each assembly step on the assembly jig, with MOSAIC setup
  Noise/occupancy
  Threshold
  Dead pixels …
Also after each layer assembled
  Repeat tests from #1, MOSAIC
After the half-detector assembled
  Final connectors & cables installed
  Need the full readout for testing?
  MOSAIC requires a special data cable/connector @PP3 (Plan-A)
    Multi-MOSAIC setups, 4?
    Test again @BNL later also
Test setup at LBNL:
MOSAIC system
  Individual Stave QA
  Good for step #1,2, and #3 with a special data cable
Small scale MVTX readout (plan-B)
  FELIX + RU + PU + mGTM
  use the final readout cables
  Multiple stave readout
  QA software and online monitoring
    Zhaozhong? Working on ITS QA,
    Mario?
@PP3: power and signal connectors
Plan-B for half-detector QA:
Connectors @PP3 are easy to manage for the final readout setup
Challenge is the “available time” to setup full readout chain @LBNL
  FELIX, RU, PU, Server
  Software development could be most time consuming
- Test each stave on the assembly jig w/ MOSAIC

## Slide 27: MVTX QA and Online Software
Can we reuse/adapt ITS2 software?
  Detector control (Win-CC required)
  Detector monitoring (No Ein-CC)
WinCC required for online control, page 3
  But not for page 4,5 monitoring outputs
Things special for sPHENIX
  New data format in progress
  FLX fdaq could be used for initial work
  sPHENIX decoder
sPHENIX software used for MVTX
Test beam data taking and monitoring
RCDAQ, with old MVTX data format
PMonitor for online monitoring