## Slide 1: IB Flexible Printed Circuit
ALICE ITS Upgrade
Stave Production Readiness Review
CERN, 27 April 2017

## Slide 2: Design requirements
Low material budget
The choice of the material to be used for the metal layers of the FPC is dictated by the need to minimise the material budget, thus Al has been preferred to the standard Cu (the respective radiation lengths being 8.9 cm and 1.44 cm).
Small losses and minimum impedance for homogenous power supply along the detector modules (9 pixel sensors row) and high quality of control and readout signals
The thickness of each Al layer is 25 μm, a value which allows keeping at the required value the voltage drop along the power supply lines; in addition, with a thickness of 75 μm for the polyimide substrate (Upilex-75S), a differential impedance of 100 Ω in the signal lines is obtained.
Dimensional specifications
The geometry of vias and interconnection pads has been optimized wrt the preliminary layout presented at the EDR, in order to allow reliable wire bonding using 3 wires per connection.
Position tolerance of ± 40 mm is required to align vias on sensor interconnection pads by overlapping the FPC on the pixel chips using nominal positions provided by alignment pin-holes.

## Slide 3: IB FPC layout
506.2 mm
16 mm
41.9 mm
16.15 mm
A
A
CHIP 1
CHIP 2
CHIP 3
CHIP 4
CHIP 5
CHIP 6
CHIP 7
CHIP 8
CHIP 9
AVDD
DVDD
75 µm UPILEX
25 µm ALUMINUM
25 µm ALUMINUM
4.1 mm
5.5 mm
20 µm COVERLAY
20 µm COVERLAY
The 9 silicon chips are read out in parallel: each chip sends its data stream to the end of Stave by a dedicated differential pair, 100 mm wide. Two additional differential pairs distribute the clock and configuration signals.
Large planes are used to distribute analogue and digital power and respective ground connections.
DGND
11.15 mm
AGND
4.1 mm
Fig 2: IB FPC Layer stackup
90 µm Glue
50 µm Silicon chip
300 µm
850 µm
p
subs

## Slide 4: IB FPC layout
16 mm
16.15 mm
10 mm
12.75 mm
13 mm
1
2
1
1
2
2
2
2
2
2
2
2
2
3
1
Oblong hole : 3 mm x 2.5 mm
2
Round hole : Ø 2.7 mm
3
Round hole : Ø 2.5 mm (for precise alignment)
Alignment and handling wings
Wings are partially cut for easier final removal at the end of Stave assembly

## Slide 5: IB FPC layout
CTRL and CLK lines
High speed data lines
AVDD
DVDD
PWELL
AGND PAD
AVDD PAD
DGND PAD
DVDD PAD
Supply and GND
SUBS

## Slide 6: IB FPC layout
1200 µm
400 µm
1200 µm
400 µm
100 µm
100 µm
350 µm
100 µm
DCTRL termination resistor
100 Ω (0402)
DCLK termination resistor
100 Ω (0402)
4 x decoupling ceramic capacitor per chip
22 uF (0402)
GND via
Power via
GND and power supply vias

## Slide 7: IB FPC layout
600 µm
400 µm
400 µm
200 µm
50 µm
50 µm
975 µm
600 µm
1600 µm
Control and data vias

## Slide 8: Inspection slits
0.7 mm
2 mm
1 mm
1 mm
2 mm
1.2 mm
0.5 mm
0.5 mm
2 mm
2 mm
2 mm
Inspection slits are implemented to verify and measure underlying chips position using reference marks

## Slide 9: FPC extension for connection to electrical services
FPC
DVDD
AGND
DGND
DVDD
DGND
DVDD
AVDD
AVDD
AVDD
AGND
AGND
DGND
The connection to the service cables is achieved by a double FPC extension which is soldered to the HIC.
x-section at interconnection between FPC and extensions
signal lines (FPC)
analogue power
digital power
FPC
AVDD
AGND
AVDD
AGND
AGND
AVDD
Flex power extension
FPC
DVDD
DGND
DVDD
DVDD
DGND
DGND
Flex power extension

## Slide 10: FPC extension for connection to electrical services
Such extensions are equipped with passive components (10x 220 mF capacitors) to stabilize the analogue and digital power supplies, respectively.
This solution allows to implement a “flexible” connection to external service cables with minimized space occupancy
Power extension Flex
IB FPC
to the Service
to the Detector

## Slide 11: FPC extension for connection to electrical services
2 Cu layer flex, PI : 50 µm, Cu : 35 µm, Solder mask : 20 µm
The 2 flexes are glued together, in the middle to have a 4 layers flex.
The PWR_extension is connected to the FPC by iron soldering.

## Slide 12: IB FPC performance: power supply drop
ANSYS simulation parameters
Power supply : Analog and Digital => 1.8V
Chip area : 3cm x 1.5 cm => 4.5 cm2
Chip power : 4.5 cm2 x 60 mW => 0.27 W / chip
2/3 for digital : 180 mW / Chip
1/3 for analogue : 22 mW / Chip

## Slide 13: IB FPC performance: power supply drop
Power plane, AVDD & DVDD DC Voltage drop
ANALOGUE | Mean | Voltage drop
Chip0 | 1.810V | + 9.75 mV
Chip1 | 1.806V | + 6.25 mV
Chip2 | 1.804V | + 3.5 mV
Chip3 | 1.801V | + 1 mV
Chip4 | 1.799V | - 1 mV
Chip5 | 1.797V | - 2.75 mV
Chip6 | 1.796V | - 4 mV
Chip7 | 1.795V | - 4.75 mV
Chip8 | 1.795V | - 5.25 mV
DIGITAL | Mean | Voltage drop
Chip0 | 1.835V | + 35 mV
Chip1 | 1.821V | + 21 mV
Chip2 | 1.811V | + 11 mV
Chip3 | 1.802V | + 2 mV
Chip4 | 1.796V | - 4 mV
Chip5 | 1.791V | - 9 mV
Chip6 | 1.787V | - 13 mV
Chip7 | 1.785V | - 15 mV
Chip8 | 1.783V | - 17 mV
Ground plane, AGND & DGND DC Voltage drop

## Slide 14: IB FPC Performance: High Speed Signal
120 µm
80 µm
120 µm
100 µm
100 µm
Al 25 µm
UPILEX-75S 75 µm
Al 25 µm
Solder-mask 20 µm
Solder-mask 20 µm
IB FPC high speed differential lines length
IB FPC aluminium conductivity : 2.40E+07 s/m
S-PARAMETERS
104.39 µm
120.95 µm
24 µm
Simulation and measurement on chip 8 HS_P data line

## Slide 15: IB FPC Performance: High Speed Signal
Chart: Frequency (Ghz)
- 8 dB Limit to no equalisation
- 3 dB @ 500 Mhz
Design validated by simulation and measurement results,
no problems for 1.2 Gb/s

## Slide 16: IB FPC manufacturing procedure
step | conditions | place | validation
1 . Thermal stabilisation of Upilex-75S sheets | 2 h at 400 oC and 1E-08 mbar | CERN or Bodycote (F) | Visual inspection
2. Holes drilling by laser | 23 +/- 1 oC (Upilex-75S CTE 20 ppm) | KIRANA (IT) | Metrology (dimensional)
3. Preparation for Al coating: cleaning, light polishing in sand blasting machine | 3 passages at 1.5 bar, 0.2 m/s, pumice powder coarse FF 10-177 um | CERN | Metrology (witness samples surface roughness)
4. Al PVD coating | 12 h of vacuum P < 10E-08 before starting | HEF (F) | Peel test, R2 measurement, Al thickness on witness samples
5. LDI, etching, Ni/Au, solder mask coverlay |  | CERN | Visual inspection, electrical continuity, resistivity measurements, metrology, metal layers thickness, “bondability” test
6. SMD components mounting by hand (baseline); under investigation vapour phase reflow oven @ 160 oC | Sn-Pb solder paste, | CERN | Visual inspection
7. Cleaning | isopropyl alcohol ultrasonic bath 3’ | CERN | Visual inspection, electrical continuity, resistivity measurements
8. Storage in N2 |  | CERN | 

## Slide 17: Fiche de suivi

## Slide 18: Upilex thermal stabilisation treatment
Upilex-75S has the lowest heat shrinkage, however this treatment turned out to be necessary to minimize and control dimensional variations during the processing
Brand/ name | Heat shrinkage [%]
Dupont/ Kapton 300 HPP-ST | 0.03 (30’ @ 150 oC)
UBE/ Upilex-75S | 0.01 (2 h @ 200 oC)
Kaneka/ Apical 300 NP | 0.04 (30’ @ 150 oC) 0.08 (2 h @ 200 oC)
Saint Gobain/ Norton TH-075 | 0.05 (2 h @ 200 oC)
2 h at 400 oC and P ~ 1E-07 mbar

## Slide 19: Laser drilling
Max four sheets of 95x28 cm2 can be mounted in the PVD system performing the Al coating; the layout has been optimized in order to maximize the number of FPC/sheet (+ production shared with MFT)
Upilex sheet
Hole pattern layout
IB
MFT
1
2
3
4
5
6

## Slide 20: FPC metrology results
The holes pattern has been designed to compensate the residual shrinking observed after Al coating (~ 0.14 %) induced by cooling down of deposited Al layers
Reference values:
  Distance (D12) between first and last via of central row in each FPC
  Position (XT0, YT0) of reference hole used for alignment wrt to first via
XT0 133.645
D12 266.99
YT0 17.66
QUOTE | nominal value [mm] | Compensated value [mm]
D12 | 266.990 | 267.364
XT0 | 133.645 | 133.830
YT0 | 17.660 | 17.684

## Slide 21: FPC metrology results
after laser drilling
Avg/stdev | I+K (1,3,4,6) | I+K (2,5) | G+H (1,3,4,6) | G+H (2,5) | PZ+PX (1,3,4,6) | PZ+PX (2,5)
After laser drilling | 267.35/0.006 | 267.39/0.01 |  |  | 267.35/0.01 | 267.40/0.01
After Al coating | 267.06/0.003 | 267.12/0.007 | 267.07/0.003 | 267.13/0.012 | 267.02/0.003 | 267.087/0.01
Finished FPC |  |  |  |  | 267.08/0.015 | 267.16/0.013
FPC (2, 5) longer than FPC (1, 3, 4, 6) by ~ 60-80 um
Finished FPCs (1, 3, 4, 6) are longer than nominal by ~60 um
St. dev. ~ 15 um
Fine tune shrink factor for the two groups
D12
after Al PVD coating
FPC thickness
Typical st. dev in one FPC ~ 4 um
On a sample of 12 pieces:
avg: 167.7 um, st. dev. 2.8 um

## Slide 22: Sand blasting
manual
sand blasting
Usage of automated sand blasting required various tests to fine tune process parameters (pumice powder size, speed, pressure) in order to achieve same Ra (~ 300 nm) as with the lengthy manual treatment.

## Slide 23: Al PVD coating
acceptance criteria | Visual inspection | P lim [mbar] | Adherence (NF 2409) | Al thickness > 25 µm TOP/BOTTOM samples | R2 [mΩ ] TOP/BOTTOM samples
Average | Aspect uniformity | 8,4 .10-7 | Classe ≤ 1 | 25,9/26,3 | 11,13/11,13
MIN |  | 7,2 .10-7 |  | 25,3/25,2 | 10,8 /10,8
MAX |  | 1,1.10-6 |  | 26,9/27,1 | 11,3/11,3
Monitoring of main process parameters (some not available to customers)
Observed values during pre-series production

## Slide 24: Metal layers thickness
FPC samples cut from not used area (the numbered rectangles) are embedded in resin and microsection photos are used to measure the thickness
Chart: SRT - Aluminum thickness (µm) - topside
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
1 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 24 | 24 | 21 | 20.5 | 2
2 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 27 | 24.3 | 22 | 23 | 24 | 18
3 | 24 | 23.66 | 23.6 | 23.6 | 23 | 22.5 | 24.275 | 25.6 | 25.2 | 24.2 | 23 | 22 | 21 | 18
4 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 23.5 | 24 | 24 | 23 | 22 | 20 | 16.5
5 | 23.5 | 23.5 | 23.5 | 23.5 | 23.5 | 23.4 | 23.5 | 23.4 | 23.8 | 23 | 23 | 22.5 | 22 | 21
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
1 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 24 | 24 | 21 | 20.5 | 2
2 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 27 | 24.3 | 22 | 23 | 24 | 18
3 | 24 | 23.66 | 23.6 | 23.6 | 23 | 22.5 | 24.275 | 25.6 | 25.2 | 24.2 | 23 | 22 | 21 | 18
4 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 23.5 | 24 | 24 | 23 | 22 | 20 | 16.5
5 | 23.5 | 23.5 | 23.5 | 23.5 | 23.5 | 23.4 | 23.5 | 23.4 | 23.8 | 23 | 23 | 22.5 | 22 | 21

## Slide 25: Resistance and impedance measurements
 | PX-001 (Ω) | PX-002 (Ω) | PX-003 (Ω) | PX-004 (Ω) | PX-005 (Ω) | PX-006 (Ω) | Length (cm)
DCLK_N | 6.624 | 6.826 | 6.331 | 6.757 | 6.707 | 6.634 | 51.27247
DCLK_P | 6.6 | 6.783 | 6.255 | 6.69 | 6.661 | 6.589 | 51.23493
DCTRL_N | 6.762 | 6.917 | 6.423 | 6.818 | 6.78 | 6.71 | 51.4434
DCTRL_P | 6.91 | 7.084 | 6.61 | 7.034 | 6.981 | 8.313 | 51.19231
HSDATA_N1 | 3.496 | 3.274 | 3.065 | 3.493 | 200 | 3.343 | 26.41708
HSDATA_N2 | 3.827 | 3.633 | 3.418 | 3.853 | 3.607 | 3.684 | 29.21161
HSDATA_N3 | 4.16 | 3.997 | 3.849 | 4.206 | 3.989 | 4.043 | 32.03476
HSDATA_N4 | 4.315 | 4.248 | 4.185 | 4.384 | 4.21 | 4.254 | 33.92726
HSDATA_N5 | 4.66 | 4.657 | 4.6 | 4.767 | 4.596 | 4.632 | 36.95352
HSDATA_N6 | 5 | 5.059 | 4.952 | 5.137 | 4.98 | 5.006 | 40.00215
HSDATA_N7 | 5.353 | 5.481 | 5.368 | 5.501 | 5.373 | 5.379 | 43.06259
HSDATA_N8 | 5.717 | 5.922 | 5.72 | 5.89 | 5.767 | 5.754 | 46.13386
HSDATA_N9 | 6.073 | 6.425 | 6.144 | 6.276 | 6.197 | 6.106 | 49.20237
HSDATA_P1 | 3.505 | 3.268 | 3.066 | 3.51 | 3.268 | 3.352 | 26.44469
HSDATA_P2 | 3.841 | 3.626 | 3.431 | 3.862 | 200 | 3.692 | 29.23948
HSDATA_P3 | 4.155 | 3.99 | 3.855 | 4.194 | 3.607 | 4.036 | 32.06206
HSDATA_P4 | 4.323 | 4.237 | 4.187 | 4.389 | 4.216 | 4.261 | 33.8915
HSDATA_P5 | 4.663 | 4.648 | 4.6 | 4.77 | 4.601 | 4.631 | 36.91777
HSDATA_P6 | 5.01 | 5.06 | 4.942 | 5.142 | 4.983 | 5.008 | 39.9664
HSDATA_P7 | 5.36 | 5.486 | 5.37 | 5.516 | 5.369 | 5.388 | 43.02763
HSDATA_P8 | 5.717 | 5.923 | 5.715 | 5.897 | 5.775 | 5.761 | 46.09889
HSDATA_P9 | 6.065 | 6.416 | 6.139 | 6.275 | 6.192 | 6.099 | 49.16661
Measured resistance HS lines on FPC PX-001 to PX-006
Chart: Time (* 2,5ps)
The curve slope is dues to the resistive loss
Probe and FPC interconnection
PX-003 HSDATA_P9 and HSDATA_N9 impedance

## Slide 26: Prototypes production
Layout and processing have been optimized during all 2016 ( more than 40 prototypes produced)
In 2017 we procured 24 prototypes, yield ~ 80 %
Al top layer after etching

## Slide 27: Production schedule
Amount to be produced: two full IB (96) + 20% → 116 FPC
Including a yield of 80% for FPC production and of 90% for HIC modules production a total number of 160 FPC is required
Each sheet contains 6 FPC (possibility to include a 7th under investigation), a total of 24 sheets need to be processed
The estimated time for full production (started in March) is 9 months, completion of 48 FPC (for IB-1) expected for mid August

## Slide 28: Production schedule